## Slide 1: AOSC 634 Air Sampling and Analysis
Practical Advice
Measuring Devices: Humidity
See Chapter 4.2 of Brock et al.
21 Feb. OK for NIST?
Copyright Brock et al. 1984; Dickerson 2015

## Slide 2: Humidity Measurement Overview
Humidity is the only routine, atmospheric chemistry measurement.
It is the most difficult of basic meteorological measurements.
Varies from 3% at the surface in the tropics to 5 ppm in the stratosphere, a range of 104.
There are no simple inexpensive calibration standards – thus the exercise with mass flow controllers.
Meteorologists think of humidity as a property of the atmosphere while chemists think of it as a component of the atmosphere.
Standard rawinsondes are only reliable to about 400 hPa.

## Slide 3: Humidity Measurement Principles
Equilibrium Sorption
  Salts or solid state sensors such as the Humicap used in the Vaisala and Davis and Hygrister used by many rawinsondes.
Extinction of Light
  Lyman alpha (VUV) or IR, fast but relative.
  Cavity Ringdown (IR) Picarro.
Condensation
  The dew pointer. There are cryogenic dewpointers with liquid nitrogen for upper air work.
Evaporation.
  Wet bulb thermometer and psychrometer.

## Slide 4
According to manufacturer

## Slide 5
http://www.vaisala.com/en/products/humidity/Pages/hmp155.aspx
Also used on Davis Met station

## Slide 6
As used on most rawinsondes

## Slide 7
Extinction of Light

## Slide 10
No dispersive element (NDIR)

## Slide 14
Developed at NCAR

## Slide 15
Lyman alpha photo-fragmentation fluorescence used on WB57 or ER2 for stratospheric water vapor (Weinstock et al., RSI, 1994; JGR 2009)

## Slide 16
Condensation
Second order systems

## Slide 20: Dewpointer pointers.
What might happen if the mirror is super clean?
Too dirty – say covered with salt?

## Slide 21
NCAR field stations
Integrated Surface Flux System (ISFS)
Deployable in networks of up to 100 stations.
State variables and fluxes.
Satellite link for DA and communication
http://www.eol.ucar.edu/instrumentation/sounding

## Slide 22: Wet bulb temperature and psychrometers
Simple, cheap, accurate, reliable
Slow, hard to link to data acquisition systems unless used with a wick and PRT.
For theory and derivation of humidity from wet bulb temperature see: http://www.atmos.umd.edu/~russ/syllabus620.html
Lecture 5A, Saturated Temp Dfn's 2012

## Slide 23: Summary

## Slide 24: Mass Flow Controller Schematics
Figure 1. Diagram of MFC. Image Courtesy of Chromatography Online.

## Slide 25: Bubbler with sintered frit for humidifying a gas flow.
Figure 1. Diagram of MFC. Image Courtesy of Chromatography Online.

## Slide 26: Schematic of selectable RH generation system
MFCd
MFCw
RH =
100%*Fw/(Fw+Fd)