AOSC 634

Experiment #7

# Ozone Measurement

Potentially Useful Equipment:  UV Absorption O3 Analyzer - Thermo Environmental Model 49, mass flow controllers, compressed air, humidity or dewpoint sensor, and chart data acquisition system (DAS).

Objectives:

1.  Determine the noise level, whether the noise is Gaussian, the effect of varying the signal averaging time on signal-to-noise level, detection limit, and response time.

2.  Compare the calibration, response times, noise levels, sensitivity to humidity, and linearity of the instrument, and generate calibration curves.

3.  Measure the amount of O3 in the ambient air over the course of a day, and discuss the diurnal variation in terms of boundary layer development, ventilation, photochemistry and dry deposition.

Procedure:

1.  Read the instruction manuals, and complete the qualifying exam.

2.  Turn on the Thermo Environmental O3 instrument with the pump off, and connect the outputs to a chart recorder or suitable data acquisition system.  Let the instrument warm up until the signal is stable, about 10 min.  Set the range to the most sensitive level, set the time response to the shortest time (4 or 10 s), and adjust the zero.  Verify that your DAS is on an appropriate setting.

3. Calculate the standard deviation of the O3 readings, create a histogram, and check for fit to a normal or Gaussian distribution.  Check and record the temperature and pressure.  Compute a detection limit.

4. Turn on the pump and let the system equilibrate about 5 min.  Check for leaks.

5. Turn on the calibrator and connect it to the compressed air cylinder.  Set the internal source for an ozone level of about 500 ppb.  Be careful to vent excess ozone out the window or into the fume hood.

6. Measure the O3 mixing ratio in the compressed air cylinder and in the output of the ozone source.

7. Make sufficient measurements of ozone at lower concentrations to determine linearity and draw a calibration curve.

8. Try changing the RH of the air sampled – this can be done with a dryer such as Nafion or by adding moisture with a bubbler.

9. Monitor ambient ozone for a day or longer; keep an eye on the weather, it may affect your results.

Hints:  Please use a filter for particles when sampling ambient air.  Remember that ozone is a reactive gas and is destroyed by contact with many surfaces, especially rubber and rubber-like plastics.

For the Report:  Calculate the detection limit for a specific signal-to-noise ratio, confidence interval (2σ level), and integration time.  Can you improve the detection limit with greater averaging times?   Do you think there is any limit to the averaging time that can be used?  From the literature, find the "clean air background" O3 level, and USEPA ambient air quality standard for ozone.  Did you experience such concentrations in College Park?   With what uncertainty can these instruments measure such mixing ratios?  Does either instrument seem sensitive to RH or rapid changes in RH?